## Sheet1
MCU Subsystem | Validation Criteria | Definition | Progress
T1 | Radar Configuration | Radar will be succesfully configured 100% of the time. | Pass
T2 | Zero Read Dropped Frames | Be able to read in data from the radar with 0% frame drop | Pass
T3 | Frame Selection | Drop any bad or incomplete frames from read with an accuracy of 100% | Pass
T4 | App Connection | Be able to send out a JSON type alert to firebase 100% of the time. | In Progress
T5 | Sliding Queue Dropped Frames | The sliding queue will be able to continually take in new frames and drop old frames without missing frames between and not failing to pass frames to the ML subsystem | In Progress
Machine Learning Subsystem | Validation Criteria | Definition | Progress
T1 | False Positive Rate | A non-fall event which is classified as a fall event as a ratio of all non-fall cases | 2.99%
T2 | False Negative Rate | A fall event which is classified as a non-fall event as a ratio of all fall cases | 2.56%
T3 | Train & Test Dataset Size | The combined number of datasets used for model training and datasets for model simulation | 164
T4 | Fall Positive Dataset | Number of datasets which serve as ground truth for falling scenario or '1' | 60
T5 | Fall Negative Dataset | Number of datasets which serve as ground truth for non-falling scenario | 104
T6 | Edge case handling | Model must account for motion patterns that do not conform to majority of fall or non-fall patterns | 20
T7 | Confuser Rejection | Combined total of false positive and false negative classifications as a ratio of total test cases | 2.83%
T8 | Noise handling | Model must account for noise that may be present in real-life residences | Added Gaussian noise of magnitude 0.1
T9 | Detection delay | Model must be able to completely process any falling motion data within a reliable timeframe | TBD in integration
T10 | Computation Speed | Real-time data must be converted to .json, preprocessed, and classified before the processing of the following data batch. A standard data batch will be 160 frames, so a dataset of this length will be used to set the benchmark. | 0.313 s
Mobile App Subsystem | Validation Criteria | Definition | Progress
T1 | Firebase Auth login | Valid users authenticate successfully; unauthenticated users are blocked from app content. | Pass
T2 | Firestore alert publish | Raspberry Pi writes alert document to production alerts collection using defined schema. | In Progress
T3 | Firestore trigger execution | Cloud Function executes on alert creation and logs successful invocation. | Pass
T4 | FCM push delivery | Push notification delivered to device for each new alert document. | Pass
T5 | Active app alert update | New alerts appear in Unacknowledged list in real time without manual refresh. | Pass
T6 | Background / killed app | Notification delivered; tapping notification opens correct alert detail screen. | In Progress
T7 | Duplicate prevention | Each alert document generates exactly one notification (no duplicates). | In Progress
T8 | User acknowledgment flow | App updates acknowledged, acknowledgedBy, and acknowledgedAt fields correctly in Firestore. | Pass
T9 | End-to-end timing | Alert reaches app within ≤ 2 seconds typical, ≤ 5 seconds worst-case. | In Progress
T10 | Burst alert handling | System handles ≥50 alerts without missed notifications or crashes. | In Progress
Power Supply Subsystem | Validation Criteria | Definition | Progress
T1 | Output Voltage within range | Output voltage should be within +- 0.1V of the expected output with both power supplied from the batteries and the AC/DC converter
T2 | Max current load achieved | Output current when connected to a load resistor is able to achive 5A
T3 | Minimum Runntime Achived | A minimum runntime of 4 hours when not charging achived when integrated with the main system achived
T4 | Battery charging voltage | Battery Charger voltage should be within +- 0.1V of the expected output
T5 | Max battery charging current | A battery charging current of 2A achived
T6 | Stable charger disconent | When the device is unplugged a spike in DC voltage of no more than +- 0.5V is seen

## New VP based on FSR
FSR Section | Requirement | Description | Margin | Status | Owner
3.2.1.1 | Responsiveness of system | The system will be able to perform detection and alert the user within a reliable timeframe of the occurrence of the fall injury event. The algorithm execution time, in addition to network latency, must remain under limit in all cases to ensure the safety of the user. This metric is defined as the time necessary to collect raw motion data, preprocess the data into the form suitable for machine learning, conduct inferencing based on said data, publish data to the cloud database, and notify the user via mobile app. | <10 s | Satisfactory | JJ Chirat William Jachimiak Preston Bowyer
3.2.1.2 | False positive rate | The system must have a low false positive rate in order to minimize the additional consumption of resources caused by false alarms, and to ensure that true positive predictions are proportionally reliable. | <5% | IN PROGRESS | JJ Chirat
3.2.1.3 | False negative rate | The system must have a very low false negative rate in order to minimize liability for the user and ensure the reliability of the system for persistent, daily usage | <5% | IN PROGRESS | JJ Chirat
3.2.1.4 | Battery life duration | The power supply subsystem must be able to provide the Raspberry Pi and radar board with battery life for reliable daily operation at full recharge. | >10 hrs | UNTESTED | Aiden George
3.2.2.2 | Mass | The mass of the full system including the Raspberry Pi, radar board, radar housing unit, custom PCB (including mass of batteries) must be within safe limits for a wall mounted installation. | <4 kg | UNTESTED | William Jachimiak
3.2.2.3 | Volume envelope | The dimensions of the full system must be within a reasonable limit such that it does not obstruct the living spaces of the user or otherwise prove difficult to handle. The acceptable limit for this system is less than 7 inches in height, 6 inches in width, and 3 inches in length. | <126 cu in | UNTESTED | William Jachimiak
3.2.2.4 | Mounting & installation | The device must be designed for easy installation in any non-hazardous location for electrical equipment. The product mount must support installation on any part of a wall. | X | UNTESTED | William Jachimiak Aiden George
3.2.2.5 | Input safety | Adherence to the recommended inputs for power & data transfer should not result in hardware damage to any component in the system. | X | Satisfactory | Aiden George
3.2.2.5.1 | Power consumption | The maximum power consumption of the system must be at or below a reasonable limit to ensure energy efficiency for the task and the safe operation of individual components | <3 W | UNTESTED | Aiden George
3.2.2.5.2 | Input voltage level | The input voltage of the system will be a built-in 7.2V as a result of the battery arrangement, which will be reduced using power electronics to a safe 5V rated for the equipment in the system. The power supply subsystem is required to remain within the safe range of input voltage dictated by parts manufacturers. | >4.9V, <5.25V | UNTESTED | Aiden George
3.2.2.6.1 | Output data | The alerts must be displayed on a mobile application that provides ease of use allows for acknowledgement by users | X | Satisfactory | Preston Bowyer
3.2.2.6.2 | Output diagnostic | The mobile app must be able to display data reated to the fall: Confidence score, location, time, and acknowledged (true or false). | X | Satisfactory | Preston Bowyer
3.2.2.6.5 | Historical Output | The mobile app must have a logistics page that holds records of alerts, graphs that show trends, and other data pertinet to risk assement on rooms/residents. The length of history should not deteriorate the functionallity ot responsiveness of the system | X | Satisfactory | Preston Bowyer
3.2.3.1 | Thermal operation | The system must operate with full capability in the specified temperature range to ensure consistent reliability in diverse environments. | -20°C to 60°C | IN PROGRESS | William Jachimiak
3.2.3.2 | Humidity | The system must operate normally in all humidity settings to prevent damage to external parts, and is not designed to operate in areas where the components may contact water, but may operate in adjacent areas (i.e. placement in higher locations). | 0% to 100% | IN PROGRESS | William Jachimiak
3.2.4.1.1 | Built-in test (BIT) | The system must supply a history of detection accuracy, including instances of false alarms and missed alarms, for the purpose of user evaluation of system reliability. | X | UNTESTED | William Jachimiak Preston Bowyer
3.2.4.2 | System Load Tolerance | The system may experience periods of elevated event activity due to sensor noise, environmental disturbances, or testing conditions. Maintaining functionality under high input load ensures system reliability and prevents loss of critical detection events. | X | In progress | William Jachimiak Preston Bowyer
3.2.5.1 | User Authentication | The mobile app must have strict user authentication because Detection alerts and diagnostic data constitute sensitive information and must be protected from unauthorized access to ensure user privacy, system integrity, and appropriate use of the system. | X | Satisfactory | Preston Bowyer
Comprehensive | Full-system demonstration | The fully integrated system must demonstrate the ability to capture motion data through the radar, infer any pattern of fall injury, and communicate with the user front-end application that a fall has occurred. This demonstration must evaluate all previous performance metrics previously mentioned. | X | UNTESTED | Comprehensive

## Adjusted FSR 0319
FSR Section | Requirement | Description | Margin | Status | Owner
3.2.1.1 | Responsiveness of system | The system will be able to perform detection and alert the user within a reliable timeframe of the occurrence of the fall injury event. The algorithm execution time, in addition to network latency, must remain under limit in all cases to ensure the safety of the user. This metric is defined as the time necessary to collect raw motion data, preprocess the data into the form suitable for machine learning, conduct inferencing based on said data, publish data to the cloud database, and notify the user via mobile app. | <10 s | Satisfactory | JJ Chirat William Jachimiak Preston Bowyer
3.2.1.2 | False positive rate | The system must have a low false positive rate in order to minimize the additional consumption of resources caused by false alarms, and to ensure that true positive predictions are proportionally reliable. | <5% | IN PROGRESS | JJ Chirat
3.2.1.3 | False negative rate | The system must have a very low false negative rate in order to minimize liability for the user and ensure the reliability of the system for persistent, daily usage | <5% | IN PROGRESS | JJ Chirat
3.2.1.4 | Battery life duration | The power supply subsystem must be able to provide the Raspberry Pi and radar board with battery life for reliable daily operation at full recharge. | >10 hrs | UNTESTED | Aiden George
3.2.2.2 | Mass | The mass of the full system including the Raspberry Pi, radar board, radar housing unit, custom PCB (including mass of batteries) must be within safe limits for a wall mounted installation. | <4 kg | Satisfactory | William Jachimiak
3.2.2.3 | Volume envelope | The dimensions of the full system must be within a reasonable limit such that it does not obstruct the living spaces of the user or otherwise prove difficult to handle. The acceptable limit for this system is less than 7 inches in height, 6 inches in width, and 3 inches in length. | <126 cu in | UNTESTED | William Jachimiak
3.2.2.4 | Mounting & installation | The device must be designed for easy installation in any non-hazardous location for electrical equipment. The product mount must support installation on any part of a wall. | X | IN PROGRESS | William Jachimiak Aiden George
3.2.2.5 | Input safety | Adherence to the recommended inputs for power & data transfer should not result in hardware damage to any component in the system. | X | Satisfactory | Aiden George
3.2.2.5.1 | Power consumption | The maximum power consumption of the system must be at or below a reasonable limit to ensure energy efficiency for the task and the safe operation of individual components | <3 W | UNTESTED | Aiden George
3.2.2.5.2 | Input voltage level | The input voltage of the system will be a built-in 14V as a result of the battery arrangement, which will be reduced using power electronics to a safe 5V rated for the equipment in the system. The power supply subsystem is required to remain within the safe range of input voltage dictated by parts manufacturers. | >4.9V, <5.25V | UNTESTED | Aiden George
3.2.2.6.1 | Output data | The alerts must be displayed on a mobile application that provides ease of use allows for acknowledgement by users | X | Satisfactory | Preston Bowyer
3.2.2.6.2 | Output diagnostic | The mobile app must be able to display data reated to the fall: Confidence score, location, time, and acknowledged (true or false). | X | Satisfactory | Preston Bowyer
3.2.2.6.5 | Historical Output | The mobile app must have a logistics page that holds records of alerts, graphs that show trends, and other data pertinet to risk assement on rooms/residents. The length of history should not deteriorate the functionallity ot responsiveness of the system | X | Satisfactory | Preston Bowyer
3.2.3.1 | Thermal operation | The system must operate with full capability in the specified temperature range to ensure consistent reliability in diverse environments. | -20°C to 60°C | IN PROGRESS | William Jachimiak
3.2.3.2 | Humidity | The system must operate normally in all humidity settings to prevent damage to external parts, and is not designed to operate in areas where the components may contact water, but may operate in adjacent areas (i.e. placement in higher locations). | 0% to 100% | IN PROGRESS | William Jachimiak
3.2.4.1.1 | Built-in test (BIT) | The system must supply a history of detection accuracy, including instances of false alarms and missed alarms, for the purpose of user evaluation of system reliability. | X | UNTESTED | William Jachimiak Preston Bowyer
3.2.4.2 | System Load Tolerance | The system may experience periods of elevated event activity due to sensor noise, environmental disturbances, or testing conditions. Maintaining functionality under high input load ensures system reliability and prevents loss of critical detection events. | X | In progress | William Jachimiak Preston Bowyer
3.2.5.1 | User Authentication | The mobile app must have strict user authentication because detection alerts and diagnostic data constitute sensitive information and must be protected from unauthorized access to ensure user privacy, system integrity, and appropriate use of the system. | X | Satisfactory | Preston Bowyer
Comprehensive | Full-system demonstration | The fully integrated system must demonstrate the ability to capture motion data through the radar, infer any pattern of fall injury, and communicate with the user front-end application that a fall has occurred. This demonstration must evaluate all previous performance metrics previously mentioned. | X | IN PROGRESS | Comprehensive

## Adjusted FSR 0330
FSR Section | Requirement | Description | Margin | Status | Owner
3.2.1.1 | Responsiveness of system | The system will be able to perform detection and alert the user within a reliable timeframe of the occurrence of the fall injury event. The algorithm execution time, in addition to network latency, must remain under limit in all cases to ensure the safety of the user. This metric is defined as the time necessary to collect raw motion data, preprocess the data into the form suitable for machine learning, conduct inferencing based on said data, publish data to the cloud database, and notify the user via mobile app. | <20 s | Satisfactory | JJ Chirat William Jachimiak Preston Bowyer
3.2.1.2 | False positive rate | The system must have a low false positive rate in order to minimize the additional consumption of resources caused by false alarms, and to ensure that true positive predictions are proportionally reliable. | <5% | IN PROGRESS | JJ Chirat William Jachimiak Preston Bowyer
3.2.1.3 | False negative rate | The system must have a very low false negative rate in order to minimize liability for the user and ensure the reliability of the system for persistent, daily usage | <5% | IN PROGRESS | JJ Chirat William Jachimiak Preston Bowyer
3.2.1.4 | Battery life duration | The power supply subsystem must be able to provide the Raspberry Pi and radar board with battery life for reliable daily operation at full recharge. | >4 hrs | Satisfactory | Aiden George
3.2.1.5* | Maximum range | The system must be able to detect falls at and below the distance specified to ensure proper coverage in the deployed space. | 7 m | Satisfactory | JJ Chirat William Jachimiak
3.2.2.2 | Mass | The mass of the full system including the Raspberry Pi, radar board, radar housing unit, custom PCB (including mass of batteries) must be within safe limits for a wall mounted installation. | <4 kg | Satisfactory | William Jachimiak
3.2.2.3 | Volume envelope | The dimensions of the full system must be within a reasonable limit such that it does not obstruct the living spaces of the user or otherwise prove difficult to handle. The acceptable limit for this system is less than 4 inches in height, 7 inches in width, and 9 inches in length. | <252 cu in | Satisfactory | William Jachimiak
3.2.2.4 | Mounting & installation | The device must be designed for easy installation in any non-hazardous location for electrical equipment. The product mount must support installation on any part of a wall. | X | Satisfactory | William Jachimiak Aiden George
3.2.2.5 | Input safety | Adherence to the recommended inputs for power & data transfer should not result in hardware damage to any component in the system. | X | Satisfactory | Aiden George
3.2.2.5.1 | Power consumption | The maximum power consumption of the system must be at or below a reasonable limit to ensure energy efficiency for the task and the safe operation of individual components | <50 W | Satisfactory | Aiden George
3.2.2.5.2 | Input voltage level | The input voltage of the system will be a built-in 14V as a result of the battery arrangement, which will be reduced using power electronics to a safe 5V rated for the equipment in the system. The power supply subsystem is required to remain within the safe range of input voltage dictated by parts manufacturers. | >4.9V, <5.25V | Satisfactory | Aiden George
3.2.2.5.3 | Battery Charging Current | The battery charging IC must be able to charge the battery at 4A and safely stop battery charging at 102% of the battery pack ideal voltage. The charging current must remain within the safe range of input current dictated by battery manufactures | >3.8A, <4.2A | Satisfactory | Aiden George
3.2.2.6.1 | Output data | The alerts must be displayed on a mobile application that provides ease of use allows for acknowledgement by users | X | Satisfactory | Preston Bowyer
3.2.2.6.2 | Output diagnostic | The mobile app must be able to display data reated to the fall: Confidence score, location, time, and acknowledged (true or false). | X | Satisfactory | Preston Bowyer
3.2.2.6.5 | Historical Output | The mobile app must have a logistics page that holds records of alerts, graphs that show trends, and other data pertinet to risk assement on rooms/residents. The length of history should not deteriorate the functionallity or responsiveness of the system | X | Satisfactory | Preston Bowyer
3.2.3.1 | Thermal operation | The system must operate with full capability in the specified temperature range to ensure consistent reliability in diverse environments. | -20°C to 60°C | Satisfactory | William Jachimiak
3.2.3.2 | Humidity | The system must operate normally in all humidity settings to prevent damage to external parts, and is not designed to operate in areas where the components may contact water, but may operate in adjacent areas (i.e. placement in higher locations). | 0% to 100% | Satisfactory | William Jachimiak
3.2.4.1.1 | Built-in test (BIT) | The system must supply a history of detection accuracy, including instances of false alarms and missed alarms, for the purpose of user evaluation of system reliability. | X | IN PROGRESS | William Jachimiak Preston Bowyer
3.2.4.2 | System Load Tolerance | The system may experience periods of elevated event activity due to various reasons. Maintaining functionality under high input load ensures system reliability and prevents loss of critical detection events. | X | IN PROGRESS | JJ Chirat William Jachimiak Preston Bowyer
3.2.5.1 | User Authentication | The mobile app must have strict user authentication because detection alerts and diagnostic data constitute sensitive information and must be protected from unauthorized access to ensure user privacy, system integrity, and appropriate use of the system. | X | Satisfactory | Preston Bowyer
Comprehensive | Full-system demonstration | The fully integrated system must demonstrate the ability to capture motion data through the radar, infer any pattern of fall injury, and communicate with the user front-end application that a fall has occurred. This demonstration must evaluate all previous performance metrics previously mentioned. | X | Satisfactory | Comprehensive